제품규격서

품 명 : 휴대용 에어로졸 측정장치

[제품사양]

□ 내역

A. 특 징

1. 휴대용 미세먼지측정기는 0.25～32㎛ 입경을 31개 채널 이상 동시에 실시간 측정할 수 있는  정밀기기여야 합니다.

- 31 Size Channel을 통해 실시간 Particle Counter 측정  (0,25-0,28-0,3-0,35-0,4-0,45-0,5-0,58-0,65-0,7-0,8-1.0-1,3-3,5-4-5-6,5-7,5-8,5-10-15-17,5-20-25-30-32 ㎛)

2. 휴대용 미세먼지측정기(광산란법 및 Filter중량법 동시측정)로서 휴대(이동) 및 설치가 용이하고, 내장밧데리를 포함하여 전원공급없이 8시간이상 실시간 측정할 수 있는 장비여야 합니다.

3. 측정기는 미국 EPA 및 유럽 EN12341의 규정에 따른 PM10, PM2.5, (PM1.0)을 측정할 수 있어야 합니다.

4. 측정기는 대기 중의 미세먼지(PM10, PM2.5, PM1.0)의 농도를 동시에 실시간으로 자동연속 측정하여 저장할 수 있어야 합니다.

5. 측정기는 저장된 데이터를 RS-232, 모뎀 또는 무선통신으로 전송할 수 있어야 합니다.

6. 측정기에 내장된 데이터로거는 대기 중의 미세먼지농도의 값을 1초, 6초, 1분, 10분, 30분단위의 데이터를 최대 1개월 이상 저장할 수 있어야 합니다.

7. 측정기는 자동 Calibration 및 Zero Span 교정확인이 가능하여야 합니다.

8. 측정기는 Hyper-terminal 및 자체소프트웨어에서 구동되어야하며, Microsoft Excel과 호환되어야 합니다.

9. 측정기는 Isokinetic Sampler set(옵션)를 이용하여 덕트 및 후드에서의 먼지측정을 환경공정시험법 중 등속흡인법에 맞추어 실시간 먼지농도를 측정할 수 있어야 합니다.

10. 측정기는 사용자가 데이터의 품질을 확인 할 수 있도록 측정기에서 발생한 오류의 종류를 시간표시와 함께 저장해야 합니다.

11. 측정기는 정밀한 측정을 위하여 자동교정확인이 가능해야하며 측정기에 발생한 오류는 측정된 데이터와 함께 저장해야합니다.

B. 구 성

1.  32ch. Real time Dust Monitor(PM10, PM2.5, PM1.0)

2.  Lead Storage Bettery

3.  Operation manual

4.  32-bit Software for Grimm EPA Dust Monitor 1109

C. 규 격

- Dust Spectrometer is designed to be a stand alone system for indoor and outdoor monitoring.

- The system can Results in  µg/m3 or particles/litre in 16 or 32 size channels or   simultaneously results according to EN 481 or US EPA

- DIN . 2066 Isokinetic sampling method

- BS 3924 Sodium Flame Photometer testing method

- ASHREA 52.1 Test Method

- US Fed. Stand. 209 for size selective dust 1998 US Fed. Stand. 209 for size selective dust

classification classification

- European EN 1994 European EN-2083 clean room size selective 2083 clean room

- EN 1996 EN -1822 high efficiency filter standard 1822 high efficiency filter standard

- American ASHREA 52.2 2000 American ASHREA 52.2